Signal processing in wavenumber and scanning distance domains of white-light interferometers

Constant term of linear fitted phase in wavenumber domain (CLPW) in white-light interferometers

Songjie Luo,1 Osami Sasaki,1,2 Ziyang Chen1, Samuel Choi,2 Takamasa Suzuki,2and Jixiong Pu1,*

1Fujian Provincial Key Laboratory of Light Propagation and Transformation, College of Information Science and Engineering, Huaqiao University, Xiamen, Fujian 361021, China

2Niigata University, Niigata 950-2181, Japan

*Corresponding author: jixiong@hqu.edu.cn

ABSTRACT

Keywords: White-light scanning interferometer; Spectral distribution;

Highlights

1. General equations of white-light scanning interferometer signals in wavenumber and scanning distance domains are made clear.

2. Relations between the signals in wavenumber and scanning distance domains are made clear. Also it is made clear that white-light scanning interferometer and spectrally resolved interferometer have a common signal in wavenumber domain.

3. In order to eliminate effects of dispersion and noises, a least square linear line is used as a phase distribution in wavenumber domain.

4. It is confirmed by simulations and experiments that a measurement value obtained by using the least square linear line provides exact measurement of surface profile.

1. Introduction

2. Signals of WLSI

Figure 1 shows a white-light scanning interferometer (WLSI) with a white-light source whose spectral intensity is I(), where  is wavenumber. The position of an object surface is zo, and the position z of a reference surface is scanned by a piezoelectric transducer (PZT). An interference

Fig. 1. Configuration of a white-light scanning interferometer.

signal is detected with a camera when the PZT is moving. The interference signal has the two components and one of them is constant during the scanning of z. Omitting this constant component,

the interference signal expressed as a function of the scanning position z is given by

,                       (1)

where phase d() is a dispersion phase caused by two sides of unequal length in a cubic beam splitter. Equation (1) is rewrite as:

.       (2)

FT (Fourier transform) of S(z) is derived as

.                    (3)

Since I() and d() are defined in the region of >0, the first term of Eq.(3) exists in the region of positive wavenumbers and this distribution is denoted by F+(). IFT (Inverse Fourier transform) ofis defined as

.                          (4)

Then, IFT of F+() is expressed as

,       (5)

where L=2(z-zO) and L is the optical path difference. This signal is the complex-valued interference signal (CVIS) of S(z) given by of Eq.(1). Equation (5) indicates that the CVIS is a

at

at

at

.

Fig.2

summation of many vectors expressed by(i=1,2,3,) where Z()=d()+4 (z-zo), as shown in Fig.2 (a). Figure 2 (b) shows the formation of SC(z) at z=za where the amplitude of SC(z) has a maximum value. Figure 2 (c) shows the formation of SC(z) at z=zp where the phase of SC(z) has zero value. Equation (1) can be reduced to a following simple equation by using Eq.(5):

.           (6)

The relations between the distribution of F+() and the distribution of SC(z) are investigated with figures. When d()=0 in Eq.(4), IFT ofis defined as

.                   (7)

Then, IFT of F+() is expressed as

.             (8)

When the distribution of I() is symmetric with a central wavenumber C, A(2z)=A(-2z) and (2z)=4Cz. The distribution of SC(z) in this case is shown in Fig.3, where A(2z) and (2z) are expressed by A(z) and (z) for the sake of simplicity. The peak position in the amplitude distribution of SC(z) is za=zo, and the zero phase position nearest z=za is zp=zo. The period of the wrapped phase distribution is 1/2C=C/2. When the distribution of I() is asymmetric with a weighted average wavenumber A, the equation of A(2z)=A(-2z) holds, but the asymmetric distribution produces a phase distribution in the z domain. The phase distribution satisfies the equations of (2z)=-(-2z) and (0)=0, but it is a little different from the distribution of 4Az indicated by the dotted line. The distribution of SC(z) in this case is shown in Fig.4. The period of the wrapped phase distribution is almost equal to 1/2A=A/2. In Fig.4, the relations of za=zp=zo holds as in Fig.3. When the distribution of I() is asymmetric and there is a dispersion phase d(), the distribution of SC(z) is shown in Fig.5. The amplitude and phase distributions of SC(z) are changed to B(L) and (L) from A(L) and (L), respectively, according to Eq.(4), and the relations of za=zp=zo does not hold. When a dispersion phase d() is produced by unequal length Ɩ of the two sides in a cubic beam splitter with refractive index, d() is given by

.                 (9)

In this case, za=zo+2b0Ɩ ??, and the position zp nearest to za at (L)=0 depends on nonlinear component of d().

Fig.3

Fig.4

Fig.5

3. Elimination of noises in signals of WLSI by using linear square line

When the position of the reference surface is not scanned at a constant speed or the optical difference path L changes due to mechanical vibrations of optical components, phase noise n(z) appears. In addition, additive noise is generated by electric devices. Considering these noise, the interference signal of Eq.(1) is changed to

.       (10)

F+() of this S(z) is expressed as

,                 (11)

where is produced from Fourier transform of at a fixed value of u. Distribution given by Eq.(11) is shown in Fig.6, where the second term of is not illustrated because it is small values compared to the first term, and ()=d()+n(). In order to eliminate the phase distribution of (), a least squares line of a0+(a1-4zo) for the distribution of -4zo+() is used as a phase distribution for F+(). IFT ofis expressed as

,                         (12)

where

za=zo-(a1/4) ,                                        (13)

zp=za-(a0WA/4)=zo-(a1+a0WA)/4 .                      (14)

Since the phase (z) is unwrapped as shown in Fig.6, the value a0W in Eq.(14) is in the range from -/2 to /2. Hence the value a0 in the least squares line of a0+(a1-4zo) is converted to the value a0W as.

a0W=a0-2m                                             (15)

where m is an integer.

D

a0W

Fig6

The values of a0 and a1 are changed by the noises, and their changes are denoted by a0 and a1, respectively. If the following condition is satisfied, the value of zp does not changed by noises;

a1+a0A=0 .                                   (15)

This condition is satisfied in Fig.7 where the least square line rotates about A by the noises as

a0=-d=-A[tan(+ )- tan()]=-A[(a1+a1)- a1]=-Aa1 .   (16)

Fig.7

The line of a0+a1 is a least square line for the distribution of (). Since the amplitude IF() has large values around A, phase changes in  () caused by the additive noise are small around A. This situation means that the least square line rotates almost about A by the additive noise. Hence the value Ea=a0+a1A is expected nearly equal to zero, that is, the value of zp is almost a constant value given by Eq.(14). On the other hand, the value of za changes by the additive noise as given by Eq.(13). If IF() is used a weight coefficient in weighted least square method for the calculation of a0+a1, the least square line more frequently rotates almost about A by the additive noise.

[It is not clear how the phase noise n() rotates the line of a0+a1. ]

----------------------------

##. Signals of SRI compared to WLSI

When the PZT is stationary and the detector is a spectral analyzer in Fig.1, the interferometer is a spectrally resolved interferometer (SRI) as shown in Fig. 6. The output of the spectral analyzer is the interference signal which is given by, where L=2(z-zo) is a constant value. Complex-valued interference signal of SRI is basically the same asshown in Fig.##. Since -2zo corresponds to L in SRI, the values of za and zp can be employed as measurement values in the same way as in WLSI.

Fig.##